Closed-Loop VA-ECMO Flow Control for Hemodynamic Stability: An In Silico Study

Mingfeng Cao, MS¹; Steve Keller, MD, PhD¹,²

¹Department of Biomedical Engineering, Johns Hopkins University School of Medicine, Baltimore, Maryland, USA

²Division of Pulmonary and Critical Care Medicine, Johns Hopkins Hospital, Baltimore, Maryland, USA

Background: Veno-arterial extracorporeal membrane oxygenation (VA-ECMO) is increasingly relied on to maintain perfusion for patients in circulatory failure. VA-ECMO provides circulatory support by acting to withdraw venous blood via a centrifugal pump and then passing it through a gas exchange device to return oxygenated blood directly to the arterial system. Current implementation relies on clinician-directed manual titration of pump speeds to deliver extracorporeal blood flow. Automated closed-loop control could enable broader deployment by maintaining hemodynamic stability without the need for expert supervision. We developed an automated controller targeting mean arterial pressure (MAP) and evaluated its performance under dynamic physiological challenges in silico.

Methods: A lumped-parameter cardiovascular model was developed and tuned using experimental data from Yorkshire swine (40 kg) that underwent severe hemorrhage to simulate circulatory failure. An asymmetric PID controller was designed with rapid RPM uptitration when MAP falls below target (65 mmHg) and modest decreases when MAP is adequate. Five SVR disturbances simulated common clinical events (vasopressor bolus, sedation, awakening, cooling, shivering) were introduced during resuscitation. Performance was compared between fixed-speed ECMO (3200 RPM) and closed-loop control, with primary outcome being time with MAP ≥65 mmHg.

Results: Closed-loop control maintained MAP ≥65 mmHg for 99.5% of the resuscitation phase compared to 64.3% with fixed-speed ECMO. MAP variability was reduced by 37% with closed-loop control (SD: 2.4 vs 3.8 mmHg). The controller rapidly increased pump speed in response to vasodilatory disturbances and conservatively reduced speed when MAP exceeded target, demonstrating appropriate asymmetric behavior.

Conclusions: Automated MAP-targeting VA-ECMO control significantly improved hemodynamic stability compared to fixed-speed operation under challenging hemodynamic conditions.

Figure 1. Lumped-parameter cardiovascular model with peripheral VA-ECMO. Eight-compartment circuit representing the left ventricle (LV), left atrium (LA), pulmonary veins (Pv), pulmonary arteries (Pa), right ventricle (RV), right atrium (RA), systemic arteries (Sa), and systemic veins (Sv). Blood flow through systemic (Rsys) and pulmonary (Rpul) vascular beds follows Ohm's law (Q = ΔP/R).

Figure 2. Closed-loop VA-ECMO control performance during simulated circulatory failure. (A) Mean arterial pressure (MAP) during onset of acute circulatory failure adapted from experimental swine hemorrhage model showing baseline (5 min), acute hemorrhage (40 min, 40% blood volume loss) to represent circulatory failure, and simulated ECMO resuscitation (30 min). Green shading indicates target range (65-90 mmHg); red shading indicates hypotension (<65 mmHg). Dashed vertical lines mark simulated clinical disturbances. (B) Pump speed during resuscitation phase showing fixed operation (3200 RPM, blue) versus closed-loop adaptation (red). (C) Percentage of resuscitation time with MAP ≥65 mmHg. (D) MAP variability (standard deviation) during resuscitation. Closed-loop control maintained target MAP 99.5% vs 64.3% of the time and reduced MAP variability by 37%.